M.S.

10/1/2010

# DRAFT 2

# Latch up and SEU test plan for Q3/Q4 2010

# Test devices

We plan to test new MAPS prototypes and all PXL detector electronics components, which are planned to be located in high and moderate radiation areas in STAR.

MAPS prototypes to be tested for latch up and SEU cross-sections are Mimosa 26 and a dedicated chip for latch up tests designed at IPHC. The LU test chip is composed of shift-registers that are built with digital cells with different spacing between NMOS and PMOS transistors. The latch up test results will help to assess the need for designing new, latch up resistant digital cells for use in sensors for the PXL detector.

PXL readout components that are planed to be located in the high radiation area (on ladders) include:

- FIN1108 Fairchild LVDS 8 Port Repeater

- SN74LVC126A  TI Quad Bus buffer gate

- SN65LVDS2 TI Single LVDS receiver

Components located on the MTB, in the moderate radiation area, include:

- FIN1108 Fairchild LVDS 8 Port Repeater

- SN65LVDT14 TI Interconnect Extender Chipset w/ LVDS

- SN65LVDT41 TI Interconnect Extender Chipset w/ LVDS

- AD7997 Analog Devices 8-ch, 10-bit I2C ADC

- CY7B9950 Cypress 200MHz PLL Clock Buffer

- MIC37152 MICREL 1.5A LDO Voltage Regulator

- AD626 Analog Devices single supply diff. amplifier

- AD8611 Analog Devices single supply comparator

- polyfuses

The AD8611 and AD626 both require a 5V power supply.

The MIC37152 and polyfuses should be studied with a 6 V power supply. All other components operate in the PXL readout system at 3.3 V.

Most of the above devices will be studied for latch up cross section and SEU. A small subset of these devices is not compatible with SEU tests or does not need to be studied in this context. This includes AD626, MIC37152, polyfuses, and AD7997 (corrupted data out of this ADC is not expected to cause any substantial problem in the system).

# Test beam

Testing is expected to take place at the BASE Facility at the 88” cyclotron at LBNL. Available beams include standard cocktails for E/A = 4.5, 10, 16, and 30 MeV. Depending on the cocktail, LETs from 1 to 100 MeV/(mg/cm2) and flux levels of up to 1E7 ions/cm2-sec are available. The beam can be as large as a few inches or as small as a quarter of an inch.

The 88-inch cyclotron website has changed recently, and the new table of Cocktails and ion species does not include the 30 MeV/nuc cocktail. If the lack of this cocktail is confirmed with the 88-inch cyclotron staff, this will lead to a minor modification of the test plan outlined below.

Switching between beams is fast and takes minutes. Changing cocktails takes approximately 4 hrs. The operating vacuum in the test chamber is reached in about 5 min and it is an automated operation. This gives a lot of flexibility in terms of swapping tests structures and moving between different ion types within a cocktail.

Parts are tested in a vacuum chamber, and can be remotely positioned horizontally, vertically, or rotationally (y and z axes) with the motion table. An alignment laser is available to ensure the part is in the center of the beam. Mounting hardware is readily available. 12xBNC (F-F), 2x25-pin D (F-M or M-F), 4x40-pin flat ribbon (M-M), 4x50-pin flat ribbon (M-M), 12xSMA (F-F), and 2xEthernet vacuum feedthroughs are mounted upon request. Holes are provided through the cave shielding blocks for connecting additional test equipment, with a distance of approximately 10 feet from the vacuum feedthrough to the top of the shielding block.

Based on our previous experience with latch up tests, we expect to use LET values that range from ~2 to 40 MeV cm2/mg.

The ion beams in the above mentioned range available at the 88” cyclotron are listed below in the order of increasing LET value. The following table was composed assuming that for LETs of similar value it is beneficial to use a cocktail that gives larger range in silicon. More information is available at the BASE facility website http://user88.lbl.gov/BASE.html. (new website: http://user88.lbl.gov/base-rad-effects/heavy-ions/cocktails-and-ions)

| Ion | Cocktail (MeV/nuc) | LET (0 deg) (MeV cm2/mg) | LET (60 eg) (MeV cm2/mg) | Range (um) |
| --- | --- | --- | --- | --- |
| Ne | 30 | 1.48 | 2.96 | 1040 |
| Ne | 16 | 2.39 | 4.78 | 347.9 |
| Ne | 10 | 3.49 | 6.98 | 174.6 |
| Si(1) | 16 | 4.56 | 9.12 | 274.3 |
| Ar | 30 | 4.87 | 9.74 | 578.1 |
| Cl | 16 | 6.61 | 13.22 | 233.6 |
| Ar | 16 | 7.27 | 14.54 | 255.6 |
| V | 16 | 10.9 | 21.8 | 225.8 |
| Cu | 16 | 16.53 | 33.06 | 190.3 |
| Kr | 16 | 24.98 | 49.96 | 165.4 |
| Kr | 10 | 30.23 | 60.46 | 113.1 |
| Y | 10 | 34.73 | 69.46 | 102.2 |

The larger penetration range will also facilitate testing of the PXL standard IC devices included in our testing plan. We have not managed to find detailed information about IC packaging composition, but it appears that it is reasonable to assume that the material density is approximately half of that of silicon.

To estimate the thickness of the IC packages, we have milled off the top layer of the IC package of the FIN1108 device. We have learned that the silicon chip is buried 10 mils (250 um) under the package surface. Similar steps will be repeated for other devices. If we can keep the thickness of the IC package layer above the silicon to approximately 10 mils (125 um silicon equivalent), this would allow us to explore a reasonable range of LET with the ion species listed above.

# Test setup

A test PCB has been designed to allow us to test all of our test devices at the same time. A simplified layout of the PCB that shows component locations is shown below:

The test card has the size of approximately 5”x 8.5” with all DUTs located in the central region in a circle with a 1.9” diameter. The interface to power supplies and a LU monitoring system is provided via 4 50-pin headers. This will allow us to conduct testing on all of the test devices in parallel and reduce the number of beam/ion changes required to accomplish our test plan. The card has been submitted for a quote. However, we already know that minor changes in the mounting holes locations will be needed to better accommodate the mounting setup in the test chamber.

For latch up monitoring we intend to use latch up protected power supplies implemented in the form of daughter cards that mount to the MTB. A dedicated support card has been designed that will host 10 of such daughter cards.

SEU tests will be performed by providing a CLK-like test signal to a DUT and comparing it with the outputs signal from the device. This can be easily accomplished with an FPGA based monitoring system. Such a system will provide much more scalability than our previous system based on NI DAQ 6800.

The number of I/O pins necessary to perform SEU test on all devices includes 10 differential pairs to interface to Mimosa26, 6 single for the LU test chip, and 22 differential pairs and 17 SE signals for other devices.

We plan to use the XC5VLX110-1FF1760 Xilinx5 development board, which is a part of the prototype PXL readout system. This provides us with access to an already implemented USB communication interface between a PC and the FPGA.

Initial development on this part of the test system has started and is expected to be finalized within a couple of weeks for the basic latch up monitoring. Additional week or two will be needed to fully implement an SEU monitoring system.

(A board diagram/schematic showing connectivity through the system should be added here.)

# Test schedule

We are currently considering 3 different schedule scenarios for performing these tests.

A)	mid November

We have reserved two shifts at LBNL’s 88” cyclotron on November 13-14 to accomplish our testing plan. As of October 1st, these were the first available time slots.

B)	October/November

As soon as the full testing setup is assembled, which is expected to happen in the second week of October, we intend to be on-call and ready to fill in empty time slots that may appear due to cancellations or earlier terminations of user runs at the 88” cyclotron. In this scenario, there is a possibility that latch up (and SEU) test results would be available before the end of October.

C)	A.S.A.P

The BNL’s TVDG test facility provides a more flexible beam access schedule, with multiple openings in October and November. However, neither X.M, L.G, nor M.S. can commit to traveling to BNL in October. This means that practically tests at BNL would be accomplished in early November at the earliest, which provides only a slight, but possibly not negligible, acceleration with respect to the schedule scenario A.