01/09/2009 XS, MS

MimoSTAR3 calibrations with LBL new RDO system

This is a short report on the first part of the MimoSTAR3 tests as outlined in detail in http://rnc.lbl.gov/hft/hardware/docs/mimostar3_testing_plan_2008_12_08.doc

The purpose of these calibration tests was twofold:

- confirm the noise measurements made by IPHC using their bias settings,

- validate the new RDO system hardware and firmware with a sensor with known characteristics.

We calibrated three chips (#2, 6, 7) with an 55Fe source at an ambient temperature of approximately 23 ºC (~27 ºC chip temperature).  These calibrations were performed using JTAG settings from our previous tests and those from the IPHC tests. The testing included both gain configurations (3 and 5). The results are summarized in Table 1 (page 2).

In addition, to verify the calibrations at different temperatures that were performed at IPHC, we tested chip #2 at approximately 9 and 19 ºC using the IPHC JTAG settings. These results are summarized in Table 2 (page 2).

Conclusions:

All gain and noise testing results are fully consistent with the IPHC measurements.

The part of the new RDO system that interfaces and reads out sensors with analog outputs has been fully validated. This includes hardware and firmware for ADC-FPGA communication, as well as firmware and software for system control and readout.

Remarks:

JTAG settings we used are based on the following set of parameters, with the main differences highlighted in bold font:

I4FASTBUF:		100 (IPHC)	40 (LBNL)

V4FASTBUF0:	92

V4FASTBUF1:	92

I4FASTINTBUF:	100

I4AMP:		100

I4REG1:		33

V4REG0-9:		49 (should be adjust for each chip)

V4TEST0-1:		150

I4PIX:			10 (IPHC)	33 (LBNL)

Table 1. MimoSTAR3 55Fe calibrations using IPHC and LBL settings at ambient temperature ~23 ºC (chip temperature of 27 ºC).

| JTAG settings | Chip | Gain | Channel | Noise run | Fe run | Noise mean (ADC) | Noise RMS (ADC) | Peak (ADC) | Peak sigma (ADC) | ENC (e-) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| LBNL | 1 | 3 | 0 | 100 | X | 3.84 | 0.35 | - | - | - |
| LBNL | 6 | 3 | 0 | 101 | 103 | 3.46 | 0.31 | 166 | 7.0 | 34 |
|  |  |  | 1 | 102 | 104 | 3.64 | 0.33 | 166 | 6.2 | 36 |
|  |  | 5 | 0 | 107 | 105 | 7.18 | 0.64 | 290 | 14.6 | 40 |
|  |  |  | 1 | 108 | 106 | 7.19 | 0.64 | 290 | 14.9 | 40 |
| IPHC |  | 3 | 0 | 109 | 113 | 4.22 | 0.38 | 230 | 5.9 | 30 |
|  |  |  | 1 | 110 | 114 | 4.21 | 0.38 | 230 | 5.8 | 30 |
|  |  | 5 | 0 | 111 | 115 | 6.53 | 0.61 | 399 | 9.9 | 27 |
|  |  |  | 1 | 112 | 116 | 6.52 | 0.61 | 399 | 10.1 | 27 |
| LBNL | 2 | 3 | 0 | 127 | 121 | 3.87 | 0.35 | 166 | 6.7 | 38 |
|  |  |  | 1 | 128 | 122 | 3.87 | 0.35 | 166 | 6.8 | 38 |
|  |  | 5 | 0 | 125 | 123 | 7.48 | 0.67 | 292 | 14.7 | 42 |
|  |  |  | 1 | 126 | 124 | 7.48 | 0.67 | 291 | 15.9 | 42 |
| IPHC |  | 3 | 0 | 129 | 119 | 4.13 | 0.37 | 208 | 7.6 | 32 |
|  |  |  | 1 | 130 | 120 | 4.14 | 0.37 | 208 | 8.0 | 32 |
|  |  | 5 | 0 | 131 | 117 | 6.19 | 0.58 | 361 | 10.8 | 28 |
|  |  |  | 1 | 132 | 118 | 6.17 | 0.58 | 361 | 11 | 28 |
| LBNL | 7 | 3 | 0 | 143 | 137 | 3.69 | 0.35 | 162 | 7.1 | 37 |
|  |  |  | 1 | 144 | 138 | 3.70 | 0.35 | 161 | 7.6 | 37 |
|  |  | 5 | 0 | 141 | 139 | 6.90 | 0.63 | 283 | 15.4 | 40 |
|  |  |  | 1 | 142 | 140 | 6.90 | 0.63 | 284 | 14.7 | 40 |
| IPHC |  | 3 | 0 | 145 | 133 | 4.02 | 0.38 | 202 | 8.5 | 32 |
|  |  |  | 1 | 146 | 134 | 4.01 | 0.38 | 202 | 8.5 | 32 |
|  |  | 5 | 0 | 147 | 135 | 5.72 | 0.90 | 348 | 12.3 | 27 |
|  |  |  | 1 | 148 | 136 | 5.72 | 0.91 | 348 | 12.0 | 27 |

Table 2. MimoSTAR3 55Fe calibrations at different chip temperatures.

| Temp (ºC) | Chip | Gain | Channel | Noise run | Fe run | Noise mean (ADC) | Noise RMS (ADC) | Peak (ADC) | Peak sigma (ADC) | ENC (e-) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 10 | 2 | 3 | 0 | 169 | 153 | 3.99 | 0.36 | 217 | 6.9 | 30 |
|  |  |  | 1 | 170 | 154 | 4.00 | 0.36 | 217 | 6.8 | 30 |
|  |  | 5 | 0 | 167 | 149 | 5.76 | 0.55 | 370 | 19? | 25 |
|  |  |  | 1 | 168 | 150 | 5.73 | 0.55 | 370 | 18? | 25 |
| 20 |  | 3 | 0 | 165 | 157 | 4.02 | 0.36 | 213 | 7.2 | 31 |
|  |  |  | 1 | 166 | 158 | 4.03 | 0.36 | 212 | 7.2 | 31 |
|  |  | 5 | 0 | 171 | 159 | 5.78 | 0.54 | 363 | 15.4 | 26 |
|  |  |  | 1 | 172 | 160 | 5.77 | 0.54 | 363 | 15.3 | 26 |